[image: image21.png](&% SCORPION

VISION SOFTWARE™





[image: image24.png]



Scorpion Laser system

For length measurements

Reference: ud-2004-0005-a Installation and Operations Manual - Laser Length Measurement
Date: June 10, 2004
Installation and Operations Manual
	[image: image2.jpg]SCORPION

VISION SOFTWARE™

fS
\




	TORDIVEL AS
Storgata 20 0184 Oslo Norway
Tlf +4723158700   (  Fax +4723158701 office@tordivel.no  (  www.tordivel.no


Scorpion Vision Software® is a Registered Trademark of Tordivel AS.

Table of content

31
Introduction


32
System description


42.1
Junction Box with camera and laser


42.2
Scorpion Vision Software CD


42.3
Scorpion Laser length measurement profile


43
Measurements


54
Installation of the system


54.1
Installing Scorpion Vision Software


54.1.1
How to get a valid license


54.2
Installing the Scorpion laser length measurement profile


54.3
Connecting the camera to the Scorpion system


64.4
Mounting of the Junction box


64.4.1
Field of view


64.4.2
Positioning the Junction box


64.5
Adjusting the laser


75
Calibration


75.1
Lens distortion calibration


75.2
Length scaling


86
Communication with external systems using RS 232


86.1
Controlling start and stop


86.2
Sending status


86.3
Sending measurement data


107
Operation


107.1
Starting and stopping the application


107.2
User interface


117.2.1
History


127.2.2
Curves


137.2.3
Results


137.2.4
Statistics


148
Settings


148.1
The conditions


159
Maintenance


159.1
Scorpion


159.1.1
Toolbox


169.1.2
Backup


169.1.3
Restore configuration


169.2
Camera


1710
Support


1710.1
Service order


1711
List of material




Revision history

	Revision
	Description
	Date

	A
	First version
	June 10, 2004


Introduction

The Scorpion Laser Measurement system is a general purpose length measurement system. 

In this document we use it for in-line measurements of lamella lengths. The system assumes that the other end of the object has a fixed position and direction during the measurement.

The camera is mounted directly over the part of the object being measured. The Laser light illuminates the object with a line. Since the laser is mounted in a 45 - 60 degrees angle versus the object, the laser line will shift position at the end of the object (see drawing). The image from the camera will look like the image below, where the shift in position is clearly displayed.

[image: image1.png]


[image: image21.png]
The Scorpion Vision Software will detect the two lines, select the line on top of the object and measure the position of the end point.

To scale the measurement a two-point calibration is required. To do the scaling one long and one short pre-measured object are needed. The objects are positioned under the camera and the end point’s y values are measured (Y1 is for the short and Y2 for the long one). These values are entered into Scorpion and the end point’s y values are scaled correctly.

To increase the measurement accuracy a lens distortion calibration can also be done.

1 System description

The Scorpion Laser measurement system consists of the following parts: 

[image: image3.png]



· Junction box with camera and laser

· Scorpion Vision Software CD

· Scorpion Laser length measurement Profile

In addition a computer is required for running the Scorpion vision Software.

The computer must have the following minimum requirements:

· Windows 2000 / Windows XP

· Intel Pentium II 255 MHz or higher

· 128 MB of RAM

· 10 MB free hard drive space

1.1 Junction Box with camera and laser

The junction box system consists of:

· Junction box with mounting plate and Montech profiles

· Unibrain Fire-i VGA camera

· Imatronic Laser

· FireWire cable

· Power supply

[image: image4.png][ —_—
aERy -
S i |

N |




1.2 Scorpion Vision Software CD

Scorpion Vision Software( is distributed on a CD with the following contents:

· Scorpion Vision Software

· System requirements

· Scorpion setup program

· Demonstration profiles

· Camera drivers

· Documentation 

· Support programs 

1.3 Scorpion Laser length measurement profile

The Laser length measurement profile is contained in a zip file to easily being sent with email. 
2 Measurements

The laser length system is measuring the lamella length. The following status values are given by the system:

· PASS

· Cannot Measure 

· Calibration Error

When status is PASS, the measured length is returned on RS232 and displayed on the screen. 

Installation of the system

2.1 Installing Scorpion Vision Software

Scorpion Vision Software is delivered on a CD.  See the Scorpion User Manual for installation details. 

In short you do as follows:

1. Turn on your computer. Insert the Scorpion Vision Software( CD-ROM disc into your CD-ROM drive.

2. Select Installation to start the installation process. 
3. Select Install - Scorpion Vision Software. Follow the instructions coming up.

2.1.1 How to get a valid license

Select www.scorpionvision.com on the CD-ROM to get a demonstration license. Copy the license code in the edit field when prompted. For a permanent license, contact Tordivel AS.

2.2 Installing the Scorpion laser length measurement profile

Install the Laser length measurement profile this way: 

1. Place the received zip-file in the ..Scorpion/Archive folder
2. Start Scorpion. Go to the Windows Start menu. Select Start - Programs - Tordivel Vision Solutions - Scorpion. A profile selection dialog is then shown.

3. In the dialog window, right click the mouse and select Restore. 

4. In the folder coming up, browse to the Archive directory where you saved the zip file and open it. 

5. The Profile is now in the dialog window and you can open the application by double-clicking on it. 
See the Scorpion User Manual for details.

2.3 Connecting the camera to the Scorpion system

Select Service in the main toolbar and give the password. 

Open the Service – Camera panel. Under Camera Settings edit the camera. Under Advanced you set the camera properties and activate the settings. You define the image format and number of frames per seconds. Using Firewire cameras it is wise to reduce the frame per seconds to a minimum.
[image: image5.png]Camera Settings

T~ HW igger

Grabber
Port

Camera

1

1

oK

Cancel

Advanced

Unibrain Fire: 1.2 - 135680

Help





See the Scorpion User Manual for further details.

Mounting of the Junction box

Before mounting the following considerations must clarified:

2.3.1 Field of view

The distance between the box (camera position) and the object defines the Field of View (FOV). Below you will find the FOV with different distances. Select the distance that will give you the area you need for your application.

The laser angle view is 40 degrees that coincide with the width of the Field of View.


[image: image6.wmf] 

Distance 

(cm)

 

w 

(cm)

 

h 

(cm)

 

Pixel resolution 

(mm)

 

10

 

7.7

 

            

 

 

 

5.8

 

            

 

 

 

0.12

 

                          

 

 

 

20

 

15.4

 

  

        

 

 

 

11.6

 

          

 

 

 

0.24

 

                          

 

 

 

30

 

23.1

 

          

 

 

 

17.3

 

          

 

 

 

0.36

 

                          

 

 

 

40

 

30.8

 

          

 

 

 

23.1

 

          

 

 

 

0.48

 

                          

 

 

 

50

 

38.5

 

         

 

 

 

 

28.9

 

          

 

 

 

0.60

 

                          

 

 

 

60

 

46.2

 

          

 

 

 

34.7

 

          

 

 

 

0.72

 

                          

 

 

 

70

 

53.9

 

          

 

 

 

40.5

 

          

 

 

 

0.84

 

                          

 

 

 

80

 

61.6

 

          

 

 

 

4

6.2

 

          

 

 

 

0.96

 

                          

 

 

 

90

 

69.3

 

          

 

 

 

52.0

 

          

 

 

 

1.08

 

                          

 

 

 

100

 

77.0

 

          

 

 

 

57.8

 

          

 

 

 

1.20

 

                          

 

 

 


2.3.2 Positioning the Junction box

The camera is positioned at one end in the Junction box and the laser in the other. 

The box shall be positioned as shown in the drawings:

[image: image7.wmf]Camera

Laser

Distance

between

object and

camera



[image: image8.wmf]Fixed

position of

the object

Junction Box

(Top view)

Camera

Laser

Field of

View


2.4 Adjusting the laser

The Laser light line must be positioned in the middle of the camera FOV.

Calibration

2.5 Lens distortion calibration

To increase the measurement accuracy a lens distortion calibration is done. This is performed by a tool named LensCalibration of type Calibrator. You find it under Service - Toolbox. You need service authorisation for doing lens calibration. Press the Help button in the tool window for more details.

2.6 Length scaling

To scale the measurement a two-point calibration is required. This operation is restricted and requires authorisation. Press the Settings button in the main panel and give the password. Select the Settings - Setup panel as seen below. 
One long and one short premeasured object are needed to do the scaling. 
1. Position one object under the camera
2. Press the Snapshot button

3. Press the Inspection button and you shall see a red cross at the end of the measured object

4. Open the Operation – Results panel and read the EndPoint value

5. Write the EndPoint value as Y1 in the Settings – Setup panel (below)

6. Measure the length manually and write the value in L1

7. Repeat step 1 to 6 for the other object and write the EndPoint and length values as Y2 and L2

8. Press Submit and the end point’s y values are scaled correctly
[image: image9.png]& Scorpion Vision Software - Laser Length Measurement

Start

g b &5 2

Settings  Service  Profile Image

I 570

Snapshot Inspection  Hep  Close

I W, 353, oo [ & sens G symenvo | () o

seun ||

640x430-5 bt

C:Program Fis!Tordivel ASiScorpion_200{Laserl engthiteasurementiimag

PASS
Length = 576.8 [mm]

Length SCORPION'
Calibration 5 VISTON SOFTWARE

Image Measurement
1

590 v:389 Pixelvalue 37 6/3/2004 12:27:17 P




3 Communication with external systems using RS 232

3.1 Controlling start and stop

Scorpion accepts the following commands for start and stop:

<stx>Start<etx>

<stx>Stop<etx>

The measured length is returned. 

3.2 Sending status 

The status values are PASS, Cannot Measure and Calibration Error.
When the status is PASS, the measured length is sent, else a 1 is sent.

3.3 Sending measurement data 

Under Service - Communication - RS232 you can see the measured data sent. 

[image: image10.png]gth Measur

@béng"l&QQ

Settings | Service  Profle  Image  Snapshot Inspection  He  Close

I Y B oy | (57 ot || By e A spenvo | () oo

General | schecer | actons | Toobos | Camera Commiricaton | watenanes | Advenced |

Tepip | [Rezzz | profibus

Options:
W Active
™ Send command respanse

COM port settings

Port rumber: 1

Parametres: [3600;8;pone; 1 set

Options:
I” showinmessages [~ Rawdata
% Show out messages

I wirtelogtofie:  [Rs232iog.txt
Protocol 1. <5T><Deta><ETX> =l
Tine. Sender Event -
1 eftond 11aast AN out 4807
1 eftojznd 11aaaB AN Out 5768
1 eftojznd 113k A Out 5973
@1 etojznd 11aaaIA Out 245
1 eftojznd 11343 A out 6.6
1 eftojznd 113428 AN Out 4807
1 eftojznd 113426 AN out 5768
o40:450-5 bt CiProgram Fles\Tordvel Asiscorpion_Z00\LaserLengthiess )1 /10/2004 11:34:23AM Ok e
1 etojznd 11ias2t A out 245
1 etojznd 11akE A Out 6.6
1 eftojznd 1ok AN out 4807
@)1 stiniomns 11as i o s 8

891 ¥:-86 Pixelvalue 64 6/10/2004 11:35:16 AM




What to send is defined under Settings – States. Select a state and you can define a command to be run when the state occurs. In our system the lamella length will be sent over RS232:

[image: image11.png]Settings for PASS

General

Constaints |

Cormmands |

Command

Guard

Parameters

RS232Cmd

5 fLengthiakie

Astivate
Edtor
Exeoute

Evecute Al





4 Operation

4.1 Starting and stopping the application

Manual start

· Start Scorpion

· Select Profile “Laser length measurement”

· The Scorpion application starts with the profile

Automatic start

You make a shortcut for an automatic startup by opening the Windows Explorer and navigate to Scorpion.exe. You find this file on the path ..\program files\tordivel as\scorpion…  Select the Scorpion.exe file, press the right mouse button and choose Create shortcut. In the shortcut properties you change the Target to System=.\<path to profile>.
Example:  ”C:\Program files\Tordivel AS\Scorpion\Scorpion.exe” System=.\LaserLengthMeasurement

You will in this example start the Laser Length Measurement profile that is placed on a sub folder. 

You can place the shortcut on the Desktop or under the Scorpion program group. Name the shortcut ’Scorpion – Laser Length Measurement’.
Stop
Stop the application this way: 

· Press the Close button in the main panel to terminate the Scorpion application. 

4.2 User interface

The screen picture under normal operation is shown below. The following information is available:

· Camera image

· Inspection result with indicator panels, green when ok, red when error

· Description – Web page that contains a short description of the inspection task

· History – displays the latest inspection results

· Curves – give a graphical view of measured values

· Results – show measured values of the latest inspection

· Statistics – give a periodical view of the inspection results

[image: image12.png]& Scorpion Vision Software - Laser Length Measurement

5 b S 2 Q

Settings  Service  Profle  Image  Close

I [ B3, coaen a1 sones | G ysonion | (Y e

Descrption | sty | Corves | Resuls | staistis |

Laser Length SCORPION®
Measurement 5 VISTON SOFTWARE

Scorpion Vision Software is an independent and open software tool making
computers see.

This profile demonstrates laser lamella length measurement with standard
components. The components used is a monachrome 640x480 firewire camera, and
laser.

with standard components you enjoy the flexibility of positioning the laser and the
camera at the distances and angle required by the measurement task.

Running on a standard PC Scorpion can process with scanrates up 25 scans/second.
The system can be enhanced by using multiple lines lasers and/or high resoluti
cameras.

« Press Start to start inspection.
© To enter Settings mode - passwor
© To enter Service mode - password = 911
« Using the Profile Selection Dialog you can make a copy of the this profile to
make your Scorpion Laser Measurement System.

C:Program Fies!Tordivel ASiScorpion_200{Laserl engthiteasurement|imagesi0003.bi

PASS Scorpion is a tool for configuration and operation of vision systems.

Length = 597.3 [mm]

1964 ¥:-128 Pixelvalue 65 6/3/2004 12:00:43 P





The Main buttons (upper row, below the main window title) are used to Start, Stop and Close the system. Additionally there are two buttons to lock/unlock the password protected Settings and Service panels. You can save an image by pressing Image and do a manual inspection by pressing Inspection. Save the profile by pressing Profile. 

4.2.1 History

The latest measurements are shown in the history list. A number identifies each image taken. The time and the classification result (the state) are additionally given in the list. By selecting an entry in the list, the image is displayed in the Image window to the left on the screen.

[image: image13.png]) Scorpion Vision Software

5'@5 9

Stop | Settings Service  Profie

I I [ B o (7 sots A syenv | ()

Descption Hstory | Curves | Resuls | statistis |

Inageno | Tme Comment -

gth Me.

6/3/2004 12:17:12PM PASS
63/2004 12:17:10PM PASS
63/2004 12:17:08PM PASS
63/2004 12:17:05PM PASS
63/2004 12:17:02PM PASS
63/2004 12:17:00PM PASS
63/2004 12:16:57 P PASS
63/2004 12:16:55 P PASS
63/2004 12:16:52PM PASS
63/2004 12:16:50PM PASS
63/2004 12:16:47 P PASS
63/2004 12:16:45 P PASS
63/2004 12:16:42PM PASS
63/2004 12:16:40PM PASS
63/2004 12:16:37 P PASS
63/2004 12:16:35 P PASS
63/2004 12:16:32PM PASS
63/2004 12:16:30PM PASS
63/2004 12:16:22PM PASS
63/2004 12:16:20PM PASS
63/2004 12:16:17 P PASS
63/2004 12:16:15PM PASS
63/2004 12:16:12PM PASS
[l 31704 17:16:10 PM_PASS.

Help

640:480-8 bi:__C:\Program Flles{Torclvel A5|5corpion_200iLasertengthieasurementimages|0002.b

812 V434 Pixelvalue 37 6/3/2004 12:17:16 P




If you stop the system, you can look at the pictures and eventually store interesting ones for further analysis. The image you choose is shown in the camera image to the left on the screen. 
Curves

You can choose between numbers of curves illustrating different values calculated by the system. All measured values can be graphed. Right click the mouse to see a menu and select add or remove curve. 

(Note: you can only add or remove a curve in Settings or Service mode. Select Settings in the main panel and give the password.)

If you are adding a curve, you get a list of all values available. The left column shows the image tools used in the analyses. The right column shows the values measured by the tool. Choose a value, click Ok and the value is graphed.

Right click the mouse over the curve, and you can reset this or all curves, open the settings panel (if you are authorized) or set the history length similar for all curves.

Adjustments of alarm limits and curve values can be done if you are in settings mode. Double click the curve name to see the adjustment parameters. You can adjust the curve axes and the alarm limits. 

The image below shows the lamella length.
[image: image14.png]i g B S 2 Q

Start Settings  Service  Profle  Image  Close

oy opsraton | ST sottnos | A systemiog Onhaut

Descton | Hitory Curvs | Resuls | states |

System.analysistime (5]

Length Value [mm]

640:480-8 bi: __C:\Program Flles{Torivel A5|5corpion_200iLasertengthiveas | 100

510 15 20 25 30 35 40 45 50 55 60 65 70 75 &0 85 80 95 100

Last| 4966 Mn| 4807 Max| 5973 Average| 5339 Stddev| 4562

i-801 V344 Pixelvalue 33 61072004 11:38:43 AM




[image: image15.png]Settings for Length. Value
s sellngs

Manvale vetcal ais [0~

Min value vertcal ais [0

Histor length 100 I~ Stais

Linits

¥ Noiy alaims
values excesds

Active dlamf oithelatest[5 hmGe

Norinl value I~ Show
Al ighfit I~ Active
Watning high it I~ Active

Watning low it I~ Active

Al low it I~ Active





Results

Measured values of each inspection are shown in this panel. 

You can choose which parameters to display by selecting New and choose from the list coming up. The left column shows the image tools used in the classification. The right column shows the values measured by the tool. Choose a value, click Add (then the window stays open – smart if you want to add more values) or Ok, and the value is included in the overview.

[image: image16.png]Scorpion Vision Software

g b 5 2 Q

Settings  Service  Profle  Image  Close

I R [ B3, | e sonas | G smenis | () s

Descrpton | Hitory | Curves Rl | sttt |

Laser Length Measurement

e T Vauo [t =
[Elenceort Porky 1860
[Eliengin Voo s mm

[ 1|

Help

640:480-8 bi: __C:\Program Flles{Torvel A5|5corpion_200{Lasertengthiveas

4 V1224 Pixelvalue 36 6102004 11:41:15 AM





4.2.2 Statistics

A periodical view of the inspection results is shown here. The table has one row for each classification state. The columns show the inspection result for this period, last period and the total.

By right clicking the mouse and selecting from the menu shown, you can manually reset the statistics. You can also choose to automatically save the statistics. The time and frequency for doing this is chosen under Scheduler in the Service panel.

[image: image17.png]Scorpion Vision Software

Laser Length Measurement

5 b S 2 Q

o o =B B O
Descton | Hitory | Curves | Resuts ot |
= e
b o e
2 e % w o
ez : : :
St : : :
o - e

|640x480-8 bit (C:\Program Files\Tordivel AS\Scorpion_2001LaserLengthMeas

T Pl :
%:4¥:224 Pixelvalue 36 6/10/2004 11:41:49 AM





5 Settings

To change the settings press the Settings button in the main panel and give the password. 

5.1 The conditions

Scorpion classifies inspection results in states. This is presented in the history list and in the inspection result field at the lower left of the screen. 

[image: image18.png]Scorpion Vision Software

g b S 2 LY Q

Start Settings  Service  Profle  Image  Snapshot Inspection  Hep  Close

s e = e [ R 1 I

Laser Length Measurement

seup 0o |

List of States

Name Description Condiion New
]
|Cannot Measure Fale Edt
|Calbration Error Fale
Down
Delete
Deactivate
Help

640:480-8 bl C:\Program Flles{Tordlvel A5|5corpion_200{Lasertengthiveas

2 V:226 Pixelvalue 40 6102004 11:42:28 AM





The following states are defined:

· PASS

· Cannot Measure 

· Calibration Error

By selecting the state and pressing the Edit button you see the state setup with constraints and commands to be sent when the state occurs. 
Maintenance

5.2 Scorpion

Authorized service operators can enter service mode when maintaining the system. Press the Service button in the main panel and give the password. The service panels to the right of the screen will then appear. See the Scorpion User Manual for operating details.

[image: image19.png]gth Measur

@béng"l&QQ

Settings | Service  Profle  Image  Snapshot Inspection  He  Close

T S, 83, e | (& e || By e B smonis | () e

General | scheer | Actions | Toobos | Caera | Commuricaton | Mantsnance | Advarced |

Profile | Options | Panels | INIFies | Misc

Profile Identification Explorer
Idertication —
Profile [CiiProgram Fils\Tardvel AsiScarpion_200iLaserlengthileasurem: | __ 910 5P8
System name [Scorpion Vision Software - Laser Length Measurement e

Project ston

Console
Pythan madle name [ScorpionProfiel
Profleverson [0z |7 I~ Read Only BED

640:480-8 bi.__(C:\Program Flles{Torcivel A5|5corpion_200iLasertengthiMeasurementiimag

41 ¥:227 Pivelvalue 37 6/3/2004 12:49:20 P




5.2.1 Toolbox

The image analysis in Scorpion is performed by a toolbox of user configurable tools. You find a rich variety of tools in Scorpion. These are image-processing tools in addition to mathematical and logical tools. They are grouped in five categories: Basic, Data, Geometry, Reference and Advanced tools. 

When run, a tool generates a result given as one or more values in addition to a set of graphical elements for visualization. The measuring result is used to define the measured objects state or status, which again decides if an action is to be taken.  The visualization elements are used to illustrate the Scorpion measurement in the camera image. Each element is given a colour.
Additionally to the image analyses tools, there are other tools used to further process the analyses results. Two tools of importance are the logical and Python tools. The Logic tool classifies results from a set of image analyses tools. 

[image: image20.png]640x430-3 bt

gth Measur

@béng"l&QQ

Settings | Service  Profle  Image  Snapshot Inspection  He  Close

L e [N

e | & s [T e | 8 i | e

Goneral| Schedler | Actions Toobox | caera | Commuricaton | Manksnance | Advarced |

Tool Settings

(C:{Program Fies!Tordivel ASiScorpion_200{Laserl engthiteasurementiimag

<

Status | 7 [ms] | Name Type Inage | Reference New

0o vi Externalscalar 1

00 vz Externalscalar 1

0o U Externalscalar 1

0o L2 Externalscalar 1

0.1 LensCaliration Calbrator 1

434 Locateline LineFinder 1

0.1 EndRef PaintLineReference I LensCalbration
166 LocateEnd LineFinder 1 EndRef

0.0 Endeoint ChangsointReference 1 LensCalbration
00 Pass LogicTool 1

25 Length Pythonscipt 1

0.0 CalbrationError LogicTool 1

[ 1|

Help

I~ Fikering
mage  +.

;
= IEE

i-608 V438 Pixelvalue 33

6/3/2004 1:07:31 PH1




5.2.2 Backup

It is wise from time to time to take a backup of the system. 

Interval

· After system delivery

· Before profile changes

· After a successful system modification

· Before software upgrades of:

· Windows

· Scorpion Vision Software

Procedure

Select the Service - Maintenance panel. Choose an archive folder and press the Backup button. For more details see the Scorpion User Manual.

5.2.3 Restore configuration

If Scorpion is upgraded to a newer version than the backup, you must remove the existing Scorpion Version (with the setup program or Windows Add/Remove program) before you restore the backup. Then you reinstall Scorpion with the same version as the backup. 
Interval

· Restore after computer breakdown

· Restore after unsuccessful profile changes

Procedure

See the Scorpion User Manual.

5.3 Camera

To change the camera, do as follows:
· Shutdown the computer and disconnect the camera cable

· Change camera and connect the camera cable

· Start the computer and Scorpion

· Follow the procedure in chapter Connecting the camera to the Scorpion system
See the Scorpion User Manual for further details.
Support
Tordivel AS is committed to give Scorpion users the best available support. 

Check our web site www.scorpionvision.com for support issues. Here you find product information, whitepapers and FAQ.
5.4  Service order

Contact can be made to:
	Phone
	+47 23158700

	Fax
	+47 23158701

	Email
	support@tordivel.no

	Address
	TORDIVEL AS

Storgata 20

0184 Oslo

NORWAY

	Description of 

service task
	


6 List of material

	Item
	Qty
	Description
	Type number

	1
	
	Junction Box system
	

	1.1
	1
	Elfa Junction Box with mounting plate
	90-810

	1.2
	
	Montech Profiles
	

	1.2.1
	1
	
	SLR-15-40

	1.2.2
	1
	
	DS-0-40

	1.2.3
	1
	
	EV-2-40

	1.2.4
	2
	
	SL-55-40

	1.2.5
	1
	
	KEV-40

	1.2.6
	30 cm
	
	AP-50-40

	1.3
	1
	Sundry accessories
	

	
	
	Camera system
	

	2.1
	1
	Unibrain Fire-i camera, VGA
	80-010

	2.2
	1
	4,5 m FireWire cable
	90-004

	
	
	Laser system
	

	3.1
	1
	Imatronic Laser, 5 mW, 40 deg angle
	90-550

	3.2
	1
	Power Supply, 5Volt
	

	
	
	Software
	

	4.1
	1
	Scorpion Vision Software – Scorpion Standard
	10-030

	4.2
	1
	Scorpion Laser measurement profile
	








� EMBED Visio.Drawing.6  ���








ud-2004-0005-a Installation and Operations Manual - Laser Length Measurement
Page 2 of 17

[image: image22.wmf]Mounting Box

[image: image23.wmf]Mounting Box

_1148124847.vsd

_1148287042.doc


Distance (cm)







w (cm)







h (cm)







Pixel resolution (mm)







10







7.7







            







 







5.8







            







 







0.12







                          







 







20







15.4







          







 







11.6







          







 







0.24







                          







 







30







23.1







          







 







17.3







          







 







0.36







                          







 







40







30.8







          







 







23.1







          







 







0.48







                          







 







50







38.5







          







 







28.9







          







 







0.60







                          







 







60







46.2







          







 







34.7







          







 







0.72







                          







 







70







53.9







          







 







40.5







          







 







0.84







                          







 







80







61.6







          







 







46.2







          







 







0.96







                          







 







90







69.3







          







 







52.0







          







 







1.08







                          







 







100







77.0







          







 







57.8







          







 







1.20







                          







 












_1148123328.vsd

_1145103042.vsd

